## Sheet1
| DELTA IMA-X600 | TDK CUS600M | ARTESYN CNS650
Length w fan enclosed | 8 | 6.18 | 7
Width w fan enclosed | 4 | 3.35 | 4
Height w fan | 1.6 | 1.67 | 1.57
Optional Main Output | 9.6-14.4 19.2-28.8 38.4-56 | 11.7-12.9, 18.5-20.5 23.4-25.9 27.3-30.2 31.2-34.5 35.1-38.8 46.8-51.8 | 12-13.8 24-27.6 48-55.2
Convection Cooled Power | not available | 400 | 400
Fan Cooled Power | 600 | 600 | 650
Fan cooled max temp non derated [℃ ] | 50 | 45 | 50
Min voltage non derated | 90 | 115 | 90
Line Regulation[%] | 0.25 | 0.5 | 2
Load Regulation[%] | 1 | 1 | 2
Minimum load | 0 | 0 | 0
power output overcurrent (stdby remain on) | hiccup | hiccup | constant current up to 50%Vo the hiccup
power output short circuit | hiccup | latch | hiccup
power output overvoltage latch (stdby remain on) | yes | yes | yes
overtemp latch | hiccup | latch | hiccup
remote sense compensation | 0.2 | 0.5 | 1.8
remote on/off | opto | opto | yes
power good | combined ac fail+dcok opto | combined ac fail+dcok opto | separate vingood, pwok
Green LED | power good | none | none
stdby [V/A] | 5/0.5 | 5/1.5 | 5/2
stdby over voltage | latch including power | no ovp sensing | latch including power
stdby ocp/short | hiccup along with output | hiccup along with output | hiccup along with output
Frequency [MinHz/MaxHz] | 47/63 | 47/63 | 47/63 (360-440w/higher leakage)
Max Inrush [Apk] | 20 | 50 | 50
STBY [V/A] | 5/0.5 | 5/1.5 | 5/2
Audible Noise [dB] | 38.5 @standby | not available | 38
Active current sharing | Yes | No | yes
IEC60601-1 3rd, IEC62368-1 approvals | Yes | Yes | yes
PMBUS | Yes | no | yes
IEC61000-3-2 Class A, PF compliance | yes | Yes | yes
EMI EN55032/EN55011-B Conducted and Radiated EN60601-1-2:2015 Immunity | yes | Yes | yes
standby state consumption [W] | not available | 3w | 0.5
Hold up 100%load [ms] | 20 | 14 | 17
Efficiency@230V | 94%@0.7load | 96%@0.5load | 94%@0.5load
Max parallel | 2 | 0 | 2
Vibration IEC60068-2-6 | 2.09grms | 10-55Hz 19.6m/s^2 | yes
Shock IEC60068-2-27 | 50g | <196 m/s^2 | yes
Variable fan speed | yes | no | yes
pollution degree | 2 | 2 | 2
Cap life [yr] | 10@80%load | 10@80%load | not available
Fan life [yr] | 7 | not available | not available
Extra features |  |  | fan override, fan fault
Note on D13: W-book: Motor load and highly capacitive load